## Sheet1
	Voltage		Unit
Power Consumed by the node	Volts	Current	Power	QTY	Net
	(PSU efficiency)	Amps	Watts		Watts	Notes
Front End Modules (FEM)	5	0.195	0.975	12	11.7	The FEM heat is NOT dissipated inside the Node Enclosure
Post Amplificer Modules (PAM) : External to Node Enclosure	5	0.4725	2.3625	12	28.35
FEM linear/analog power generation and distribution	0.45		14.3	1	14.3	linear power supplies are about 45% efficient
PAM linear/analog power generation and distribution	0.45		34.65	1	34.65	linear power supplies are about 45% efficient
PAM margin	5	2	10	1	10
SNAPv2	12	2	24	4	96	1.6Amp is a measured values; use 2A for margin
White Rabbit	5	1.6	8	1	8
Beaglebone	5	2	10	1	10
+12V switching/digital power generation and distribution	0.8		24	1	24	switching power supplies are about 80% efficient
+5V switching/digital power generation and distribution	0.8		4.5	1.00	4.5	switching power supplies are about 80% efficient
Interface Logic and Monitoring	5	1	5	1	5
Externally mounted centrifugal fan	230	0.32	66	1	66	The Fan heat is NOT dissipated inside the Node Enclosure. Operating point on curve is 0.32A 66W. The fan is rate for 0.36A and 90W.

Total					312.5

Power dissipated inside the Node Enclosure					234.8
Power dissipated at the feeds					11.7
Power dissipated at the centrifugal fan					66	This heat will be handled by long run of buried ducting between the centrifugal fan and the Node Enclosure

input VAC voltage	230

Input VAC currents (nominal operating mode)	power factor	VAC input Amps
FEM linear/analog power generation and distribution	0.65	0.174				generic value from generic linea power supply : http://www.acopian.com/linear-power-supply-vs-switching-power-supply-vs-unregulated-power-supply.html
PAM linear/analog power generation and distribution	0.65	0.488				generic value from generic linea power supply : http://www.acopian.com/linear-power-supply-vs-switching-power-supply-vs-unregulated-power-supply.html
+12V switching/digital power generation and distribution	0.93	0.561
+5V switching/digital power generation and distribution	0.93	0.129
Externally mounted centrifugal fan	0.90	0.320				(assuming fan is at operating point 2 (see datasheet)

VAC current input to the Node Enclosure		1.352
VAC current input to the Node (Total)		1.672				Includes current for external centrifugal fan

Inrush Input VAC currents
When the Node Enlcosure is "turned on" only the +5V supply turns on.
Then if the internal computer decides everything is ok then it will turn on the other VAC consumers and it can do this 1consumer at a time.
FEM linear/analog power generation and distribution	40					total guess for the TBD Lineare power supply inrush
PAM linear/analog power generation and distribution	40					total guess for the TBD Lineare power supply inrush
+12V switching/digital power generation and distribution	50
+5V switching/digital power generation and distribution	35
Externally mounted centrifugal fan

Air Flow	106	CFM
Air Flow	3.002	Cubic meters per minute
Pipe Diameter	110	mm
Pipe Wall thickness	3	mm

Pipe Area Interior	8495	mm squared
Pipe Area Exterior	9503	mm squared
Air Speed	353.341129760739	meters per minute
	5.88901882934565	meters per second
	17.8195921972454	ft per second

http://www.engineeringtoolbox.com/convective-heat-transfer-d_430.html
heat transfer coeff	30	Watts / (meter squared * Kelvin)

SNAP back cover surface area	80	inches squared	ignoring heat path to top cover
	0.0516128	meters squared
Max FPGA junction temperature	80	Celsius
Thermal Resistance from junction to case	0.2	Celsius/Watt
Thermal Resistance from case to PCB	0.1	Celsius/Watt
Thermal Resistance of the Gap Pad 3.0 bottom	1.5	Celsius/Watt
Thermal Resistance of the bottom heat sink	0.0495	Celsius/Watt
Thermal Resistance of Heat sink to the bottom cover	0.02	Celsius/Watt
Thermal Resistance from junction to bottom cover	1.8695	Celsius/Watt

FPGA power as function of SNAP power	0.8
Maximum allowed bottom cover temperature	44.1056	Celsius
Maximum node air temperature	28.605568999938	Celsius

Air Temperature of node inlet year round	21	Celsius
Air Temperature margin	4	Celsius
Rise in air temperature dut to PAM	3	Celsius

Maximum node air temperature